## Slide 1: Introduction to XOPPY
Basics X-ray Optics calculations

## Slide 2
Introduction to XOPPY
Transport of power in a Beamline
Manuel Sanchez del Rio
srio@esrf.eu
Advanced Analysis & Precision Unit, MEG/ISDD, ESRF
May 29th 2023
XOPPY

## Slide 3: Transport of power in a Beamline
Outline:
Basics
History
Source emission: Flux and power spectra.
Optical elements: characteristics and absorbed power.
Flux and power that arrive to the sample
XOPPY

## Slide 4: Basics

## Slide 5: The Beam: Spectrum and Size
Size: FWHM
If profile is Gaussian
FWHM = 2.355
x
y
A light beam or beam of light is a directional projection of light energy radiating from a light source.
Collection of particles [photons]
Collection of waves
The main parameters can be represented:
As histograms if they are discrete variables
As functions if they are continuous variables
Spectrum I(E) Beam profile I(x,y) [cartesian]
I(E,x,y)
CORRELATED!

## Slide 6: Intensity vs E
I can be
Flux: Photons per second per (variable) bandwidth (0.1% E)
Spectral power: Power per (constant) bandwidth W/eV
Spectral power = Flux x constante (constant = 1e3 * 1.602176634e-19)
Total power = integral of spectral power
cte = 1e3 * 1.6e-19

## Slide 7: I(E) and I(x,y) are correlated (3D dependency)
I(x,y,E):
Flux vs (x,y,E): Photons/s/0.1%bw/mm^2
Spectral power vs (x,y,E): W/eV/mm^2
Power density: integral over E: W/mm^2
Spectral power: integral over (x,y): W/eV
Total power – triple integral

## Slide 8: How I is calculated?
For synchrotron beams, considering
The classical electrodynamics (electric field from a moving charge)
The phase space volume of the moving charges (~ storage ring emittance)
It is much simpler to work with electric fields than with intensities.
=
Waves: Scalar complex amplitude (at a given position z wavefront )
Rays: Real vectors:
Example: Undulator

## Slide 9: How does the beam propagates in free space?
If rays: Need direction or angle and propagates as a straight line (without changing the intensity)
If waves we need an integral propagator: Fresnel etc. (changing the intensity distribution, conserving the integrated intensity
?
z1
If we know how to propagate, we can define Brightness or Brilliance:
peak intensity, on-axis, and far away
[photons/s/0.1%bw/mm^2/mrad^2]
z2

## Slide 10
How the beam propagates through components – Optical Elements
Geometrical model (change propagation) Physical model (change the intensity)
[change the ray direction or the wave phase] [change the intensity via the electric field]
IDEAL LENS IDEAL ATTENUATOR
material & density
thickness
X-ray beam
f
Beer-Lambert Law

## Slide 11: Transfer Matrix: the simplest ray-tracing
“linear” black box
Straight propagation: distance x divergence [not quadratic!]
[same divergence]
Lens: : [no-propagation]
[direction changes with x]
Mirror: (normal incidence) [grazing incidence, Coddington equations]

## Slide 12: Transfer Matrix: typical p:q focusing
= =
M [magnification]
[1/M + a linear term with x0]
p
q
S
s = S x M
Lens equation Magnification

## Slide 13: Other matrix-based systems: The electron beam dynamics
General
No z-coupling
Finite alpha
Symmetry point
Statistically distributed bunches
P1
P2
P3

## Slide 14: Extending Transfer Matrix: Phase Space Analysis, Ray tracing
Phase Space Analysis:
https://doi.org/10.1364/AO.47.00E116
Too basic:
Gaussian-shaped beam at the source
Gaussian elements
Linear (matrix) operators.
These limitations are not in full ray tracing [shadow]:
18D phase space
fully non-linear operators (able to calculate aberrations, surface errors, tolerances, etc.)
…

## Slide 15: Hand-calculations. Example ESRF ID16A. Undulator source
U18.3 17keV
L=1.4m
Ref. M. Sanchez del Rio et al, JSR 26 1887 (2019)
https://doi.org/10.1107/S160057751901213X
Single e-
Divergence:
Size:
Photon source
Photon single electron emission at wavelength 0.719735 A:
sigma_r: 2.19583 um
sigma_rprime: 4.93133 urad
Electron sizes:
sigma_x: 30.1836 um
sigma_z: 3.63641 um
sigma_x': 4.36821 urad
sigma_z': 1.37498 urad
Photon source sizes (convolution):
Sigma_x: 30.2634 um
Sigma_z: 4.24796 um
Sigma_x': 6.58781 urad
Sigma_z': 5.11944 urad

## Slide 16: Hand-calculations. Example ESRF ID16A Beam sizes and flux
Flux: use geometrical crop
  Initial (source): 2.9e14 / 17 ph/s/eV
  Monochromator: x 1e-2 * 17000 = 170
  Crop VSS (none for EBS)
  KBV (60mm) 7.1%
  KBH (26mm) 40%
  Total 2.85%. Thus 8.3 1013
Apply mirror reflectivity

## Slide 17: Hyerarchical simulations with Oasys
Analytical model (by hand and XOPPY)
Ray tracing (ShadowOui)
Hybrid model (ShadowOui)
Simplified wave optics (WOFRY) Coherent Mode Decomposition (2D CMD also possible)
Monte Carlo (multi e-) wave optics (SRW)
Ref. M. Sanchez del Rio et al, JSR 26 1887 (2019)
https://doi.org/10.1107/S160057751901213X
COMPLEXITY

## Slide 18: XOPPY history
XOP (X-ray Oriented Programs) (http://dx.doi.org/10.1063/1.1147376)
          a user interface to useful tools for the beamline designer
                collaboration ESRF-APS (M. Sanchez del RIO and Roger J. Dejus)
                it included some tools for synchrotron spectra (US, URGENT), and a bunch of pieces of SHADOW to compute BM and wiggler sources, crystal, multilayer and mirror reflectivities.
                The interface was written in the IDL binding Fortran codes
XOP evolved (2004-2011) with mode dedicated codes and a new optical constants database (DABAX) (http://dx.doi.org/10.1117/12.893911 )
The XOP concept migrated to OASYS, and new interfaces were created in python (therefore the name of the add-on XOPPY) but always maintained the initial philosophy of “toolbox”
OASYS-XOPPY uses xraylib. Many new features have been added: 3D calculations (E,x,y) SRCalc Scripting DABAX available
1996
2004
2009
2013-2023

## Slide 19: Sources

## Slide 20: Simulation of source emission
XOPPY
Source spectrum (Photon flux or spectral power vs Energy)
bending magnet
Wiggler
Bending
magnet
Undulator

## Slide 21: Undulator emission simulation
U42
Slit
30 m
ESRF ID20 U32c
“red shift”
Undulator period
Photon wavelength
(harmonics)
XOPPY

## Slide 22: Undulator emission simulation: Peak position vs aperture size
XOPPY
U42
Slit
30 m
~50 eV
1st harmonic

## Slide 23: Undulator emission simulation: Peak position vs K
XOPPY
U42
= 1.2398e-06
Memorandum:

## Slide 24: Undulator emission simulation: Peak position vs K
XOPPY
U42

## Slide 25: Undulator emission simulation (gap)
Gap

## Slide 26: Undulator emission simulation
Gap
Reducing
the gap
incrementing K
larger wavelength

## Slide 27: Simulation of source emission
Undulator Spectrum (Flux/Spectral power)
Undulator power density (power vs x, y)
Undulator radiation (flux/power vs x, y, E)
Wiggler Spectrum
Wiggler radiation

## Slide 28: Simulation of source emission
Undulator spectral flux density
(flux/power vs x, y, E)
Wiggler [and short ID’s] spectral flux density
(flux/power vs x, y, E)
y
E
Photon flux
Photon Energy [eV]
1000 5000 10000 15000 20000
x
Power density distribution of ESRF-3PW
https://jacow.org/medsi2020/papers/wepb17.pdf

## Slide 29: Undulator power density projected (Power vs x, y)
For example, this tool is very useful to get the heat load on a slit:
undulator
Water cooled high power slit

## Slide 30: Undulator power density (Power vs x, y)
In most of the beamlines at the ESRF the are horizontal slits at 16 m from the source:
X-ray beam
Project over the slit surface (2o angle)
Perpendicular to the beam
Finite element analysis

## Slide 31: Optical Elements

## Slide 32: Material constants libraries xraylib https://github.com/tschoonj/xraylib
Other apps with other data:
XCOM: https://physics.nist.gov/PhysRefData/Xcom/html/xcom1.html
S. Stepanov server: https://x-server.gmca.aps.anl.gov/
CXRO server: https://henke.lbl.gov/optical_constants/

## Slide 33: Power transport on a beamline
Optical components that could be present in a beamline:
Source
Front end
Water-cooled
slits
Windows
Primary slits
Primary mirrors
Monochromator
Focusing system
Sample
KB mirrors

## Slide 34: Power transport on a beamline
Optical components that could be present in a beamline:
Slit 1
Projected size at
downstream element (A)
Distance from the source (d)
Aperture (a)
Distance from the source (D)
Source
Source
Front end
Water-cooled
slits
Windows
Primary slits
Primary mirrors
Monochromator
Focusing system
Sample
KB mirrors

## Slide 35: Filter or attenuator (window) absorption
Filters of different thickness
Filters of different materials

## Slide 36: Power absorbed by a filter or attenuator
Heat load on filters @ 23 m with a projection of 2 mm x 1 mm:
U42-EBS
Material | Thickness [mm] | Absorbed Power [W]
Be | 300 | 19
Diamond | 300 | 40
thickness
material & density
X-ray beam

## Slide 37: Power absorbed by a filter
Heat load on filters @ 23 m with a projection of 2 mm x 1 mm:
Be (300 mm)
Diamond (300 mm)

## Slide 38: Power transport on a beamline
Optical components that could be present in a beamline:
Slit 1
Aperture (a)
Source
Absorption by filters
Absorption power by mirror
Source
Front end
Water-cooled
slits
Windows
Primary slits
Primary mirrors
Monochromator
Focusing system
Sample
KB mirrors

## Slide 39: Mirror absorption
Same material, different angles
Same angle, different coatings
Fresnel equations give the reflectivity as a function of angle and photon energy. As a consequence, one gets the critical angle:
Si s-reflectivity

## Slide 40: Mirror absorption
1st Harmonic
3rd Harmonic
1st Harmonic
3rd Harmonic
Example of reflection, primarily mirrors @ 30 m
Absorbed power 113 W
Absorbed power 52 W

## Slide 41: Power absorbed by a mirror
Heat load on mirrors @ 30 m:
Si (7 mrad)
Diamond (100 mm) + Rh (3 mrad)

## Slide 42: Calculation of flux at the sample position
Example of getting the flux at the sample position:
KB slits
51 m
Source
Front end
Water-cooled
slits
Primary slits
Primary mirrors
Monochromator
Focusing system
Sample
KB mirrors
*
XOPPY: Total photons per second 1.32091E+13 at 5000 eV
ShadowOui: Integrated Flux : 1.25338e+13 ph/s

## Slide 43: Other optical elements in XOPPY: Multilayers
[W(25)/Si(25)]50
1.4o

## Slide 44: Other optical elements in XOPPY: Crystals
Crystals
2q
Intensity
Laue's kinematic theory:
small crystals (e.g. powder diffraction).
Laue's kinematic theory:
small crystals (e.g. powder diffraction).
Dynamical theory of diffraction:
Large perfect crystals (e.g. Si).

## Slide 45: FORMULATIONS OF DYNAMICAL THEORY
Darwin (1914)The Theory of X-ray Reflexion Philosophical Magazine 27, 315
Ewald (1917) Zur Begründung der Kristalloptik, Teil III Die Kristalloptik der Röntgenstrahlen Ann. Phys. 54, 519
von Laue (1915) Die Theorie der Röntgenstrahlinterferenzen in neuer Form Encyclopaedie d. math Wiss 5,3, (1931) Ergebn. Exakt. Naturwiss.10133.
Zachariasen (1945) Theory of X-ray diffraction in crystals
James (1950) The optical principles of diffraction by X-rays
Taupin (1961) Théorie dynamique de la diffraction des rayons X par des cristaux déformés Bull Soc Fr. Mineral. 84 51
Takagi (1962) Dynamical theory of diffraction applicable to crystals with any kind of small distortion Acta Crystal 15 1311
Books:
Pinsker Dynamical Scattering of X rays in Crystals (1978).
Authier Dynamical Theory of X-Ray Diffraction (2001)
Shvyd’ko X-Ray Optics—High-Energy-Resolution Applications (2004)
Si111
10 keV
Diffraction profile
or
Intrinsic profile

## Slide 46: Rocking curve
The rocking curve is the CONVOLUTION of the two diffraction profiles.
For a full discussion about rocking curves see: Masiello et al. J. Appl. Crystall. 47, 1304-1314 (2014)
FWHMR = 1.4 FWHM1
=
FWHMR = 1 FWHM1
=
FWHMR ~ 1.5 FWHM1
Si111 8 keV
=

## Slide 47: Crystal phase retarders
Perfect crystals are birefringent elements: Rs  Rp
They introduce a phase change F that alter the polarization of the beam
T
R
BRAGG
R
T
LAUE

## Slide 48: Change of phase
Hirano et al., NIM A336 (1993) 343
4XGe220
50mm Si220
22 mm C* 220
22 mm C* 220
BRAGG
(REFLECTION)
LAUE
(TRANSMISSION)

## Slide 49: Circular polarization
Bouchenoire, Morris, Hase
Si111 5mm
C*111 128mm
E=3.174 keV
Pc = rs rp sin (fp - fs)
XOP